I. Hardware Setup

- Computer

- Power to white box

- Power to silver box

- Turn on computer

- Password: rogue123

- Launch Brainsight→Agree to terms

- Magstim

- Power switch on rear of all units

- Turn on all units (front panel, top left corner, black button)

- Set stimulation magnitudes (suggest 50/0/0)

- Front panel, bottom right, green dial

- Turn to adjust first value, press to move on to next value

- Calibrate Coil

  - Place black bracket on silver paddle

  - Window→TMS Coil Calibration→Alpha→Re-calibrate

  - Place coil on holder, hold still in camera field of view

  - Speak “Begin Countdown”, hold very still

  - Repeat until successful

  - Close calibration windows

- Magnet Safety

- Remove phone and wallet from researcher’s pockets

- Remove phone and wallet from participant’s pockets

- Participant Hardware

  - EMG

- Attach Velcro wristband

- Clean electrode sites with alcohol wipe

- Negative: belly of muscle (have participant flex thumb and forefinger to find belly)

- Positive: Tendon, medial side of finger between 3rd & second knuckle

- Ground: Bony knob on lateral part of wrist

- Attach electrodes

- Participants hand should be in a relaxed facing up position

  - Motion Capture

- Screw triangle set into headband (located in black suitcase)

- Put on participant’s head

- Make sure triangle is opposite to side of stimulation

II.a Software Setup: MNI template

(skip if individual T1 is available)

- New MNI Head Project

II.b Software Setup: Individual T1

(skip if individual T1 is not available)

- New Empty Project

- Anatomical

  - File: Click Here to choose…Browse for T1 (any format, unzipped)

  - Show Image & Details

- Check that file loaded ok

- Register MNI with T1

  - Click on Atlas Spaces

  - New→Manual (AC-PC-Box)

  - Navigate to Anterior Commissure→Set AC

  - Navigate to Posterior Commissure→Set PC→Next Step

  - Click & drag green box to L/R, A/P, D/V boundaries

    - Anterior/Posterior use sagittal view

    - Dorsal/Ventral use coronal view

    - Left/Right use axial view

  - Finish

  - Apply then close

Sample Registration

- If functional localizers are available

  - Use other software to make thresholded statistical maps from fMRI

  - Resample these images to same resolution as T1

  - Overlapys tab→Configure Overlay

  - Add… (top left) →Browse for function

  - Navigate to check that images are properly registered

  - Close

- Reconstruction

  - New→Skin→Compute Skin→close

  - New→Full Brain Curvilinear→Compute Curvilinear

    - Adjust peal depth (top left scale) to reveal area of interest (~4mm)

    - If fMRI localizers are available: Click blue i→Overlays→Show

  - Close

- Landmarks

  - Configure Landmarks

    - Nasion (concavity above the nose)

    - Tip of Nose

    - Back of head (bottom of inion)

    - Left & Right Ear (supratrageal notch)

  - Click each location of Skin surface→New→Give a sensible name

  - Make sure points are easy to replicate on actual head

  - Close

III. Define Targets

- Targets tab

  - Configure Targets

- Set Layout

  - Set layout to 1|2

  - Set largest panel to curvilinear

- Set Targets (M1-hand grid-search)

  - Click brain surface centered on area of interest

  - Use sliders to set orientation (suggest 20, 40, 140)

  - New… (Top left)→Rectangular Grid

    - Name = M1_10x6 (or whatever)

    - 3D size = 3

    - Kind = Trajectory

    - Grid size = 10 X 6

    - Spacing = 5mm X 5mm

  - Set Curvature

- Rotate curvilinear brain to view grid from front

- Adjust curvature slider until grid hugs surface

- Suggest about 150mm

  - Note: To nudge the center of the grid

    - click new center location→Move target to Crosshair Origin

  - Close

Sample 10 X 6 Grid over hand M1

IV. Perform

(Everything up to this point could be done before participant arrives)

- Start the session

  - Session → New Online Session (opens new window)

- Carry over Targets

  - Select from left panel

  - Click Add

  - To select all  click the top one + shift + select the bottom one

  - Click Randomize Selected

  - Repeat for number of desired samples per target

  - Next Step

- Input/output

  - Check Use TTL Channel 1

  - Check Acquire EMG Channel 1

  - Check Acquire EMG Channel 2 (if used)

  - Auto Compute

  - Ask participant to wiggle thumb slowly. Does the EMG track?

  - Next Step

  - Confirm that patient head is in camera field of view

  - Next Step

- Registration (might already be done for custom T1)

  - Turn on voice recognition

  - Place registration gun at participant’s Nasion

- Concavity between the nose and brow

  - Speak “Sample”

  - Repeat for LPA & RPA

- Concavity above ear notch

  - Repeat for inion

- Back of head

  - Next Step

- Scaling

  - Use registration gun to identify landmarks

  - Ask participant to tilt head if necessary

  - Important: Use 2D views to ensure that selected location are at the extreme edges of brain

    - Frontmost (check sagittal view)

    - Topmost (check sagittal view)

    - Backmost (check sagittal view)

    - Rightmost (check axial view)

    - Leftmost (check axial view)

Note: If validation fails, consider adding additional points over area of interest.

- Validation

  - Run registration gun across scalp slowly and firmly

    - From anterior point to posterior point

    - From left to right

    - Check particularly around target area

  - Watch Crosshairs  Skin value.

  - If all Green/Orange range, continue

  - If red, particularly in area of interest, repeat steps 6 and/or 7

  - <5mm ok, <2mm good.

  - Next Step

- File Save Project As…

  - ID#_Session#_YYYYMMDD

  - Click Resume

V. Data Collection

- Neuro-navigation setup

  - Arrange windows 2 X 2

- Curvilinear Brain & Samples

- Bullseye (Coil Centric) [Driver = Alpha]

- EMG

- EMG Pod Recording

  - Top left list  select first sample

  - Bottom left list → adjust so Error and EMG channel 1 are visible

  - Bring coil to target. Prefer errors <1mm

    - Circle = location on skin surface

    - Dot = location on brain surface

    - Line = orientation (should face up)

- Science!

  - Arm TMS machine by pressing the green button on the top unit (orange light will flash on when machine is armed)

  - Press peddle and buttons to zap

  - Speak “Next” to change targets

  - Rinse and repeat

- Leave about 3s between stimulations

- Trouble Shooting

  - Is Magstim armed?

- Should have orange indicator light, topright corner

    - If not, press right-side green button

  - Did the pedal time out?

    - Step off then back on the peddle

- File Save Project

  - Close Perform Window

VI. Post-Processing

- Review Data

  - Session→Review

  - Layout = 1X2 (curvilinear in largest screen & EMG)

  - Top left Targets window  check/uncheck Targets to show/hide grid locations

  - Underneath Targets window, see Sessions window  uncheck Session 1 to hide sample locations

- MEP heat map

  - Blue i icon (middle top)

  - Motor Maps

  - Check Show Map

  - Check Colour Samples

  - LUT=Hot

  - Threshold about 50 -1500 (depending on range of observed MEPS)

  - Change full width half max to 5mm

  - Update Selected Map

- Export Data

  - Export (bottom-left) →Brainsight Text file

  - Choose things to save, probably everything

- Name file ID_Session_YYYYMMDD

  - Save

    - Tab 	delimits columns

    - ; 	delimits rows

VII. Take Down

- Make sure everything is saved

- Return equipment

  - Return to cart EMG equipment to cart

  - Return tracking equipment to black case

- Close all windows

- Power Down Computer

  - Turn off computer

  - White box power switch

  - Silver box power switch

- Power Down Magstim

  - Power buttons (all units)

    - Front panel, left side, black button

  - Magstim power switches (rear panel)

Phase 1: Mapping

      - Stimulate across target grid

        - Use grid targets to map both hemispheres

        - Use review panel to make heat maps for both hemispheres

        - Define new targets at the peak of each heat map

        - Might be necessary to convert an observed stimulation site to target

Phase 2: Motor Threshold

      - Positioning

        - Place chinrest for participant

        - Let’s take extra care to make sure that their comfortable

        - Navigate left hemisphere coil to target and fix in place

        - Navigate right hemisphere coil to target and fix in place

        - Readjust left of necessary.

        - Participant should stay still now.

      - Code setup

        - In python find the names of the serial ports for the stimulators

        - Should be something like tty.usbserial, and tty.usbserial1)

        - In Terminal:    /dev/tty.*

        - Open MT.py and check that MagOne = …. Has the name of a serial port

      - Test some thresholds for one hemisphere

        - Warm participant that some stimulation will start

        - Navigate to folder

          - In Terminal: cd Experiments/BiHem

        - Run a stimulation train

          - In Terminal: python MT.py

        - This will ask you for a motor threshold to try (start around 45)

        - Watch Brainsight for MEPs, count number >50mV

        - If half of stims elicit MEPs keep this value

        - If not try another value

      - Repeat for other hemisphere with other serial port.

Phase 3: IHI (~15 min)

        - Leave stimulators set to motor thresholds, code adjusts automatically

        - In Terminal: python BiHem.py

        - First answer some questions about the participant

        - Look at the stimulators and see which display has #1

        - Warn the participant that stimulation will start soon

        - Answer some questions about the stimulators

        - Let this run for a while. The code might now close properly this is some problem that I haven’t solved yet. All of the data are collected ok, just close Terminal and shut off the computer.

        - In brainsight be sure to save the session and export the EMG data.